體外生命(心肺)支持裝置-Extracorporeal life support (ECLS)

### 詹 聖 霖

ECLS包括cardiopulmonary bypass (CPB), extracorporeal membrane oxygenation (ECMO), cardiopulmonary support (CPS) or extracorporeal cardiopulmonary resuscitation (ECPR), ventricular assist devices (VAD)

- Indication:

_____ 1.High mortality risk disease but most likely reversible

_____ 2.Failure to respond to maximal conventional medical treatment

Acceptable criteria for ECLS

_____ Pulmonary support:

  - ____ Inadequate oxygenation (respiratory failure): MAS, RDS, CDH, PPHN, ARDS --

PaO2/FiO2<100-150 or AaDO2>450-500mmHg or SatO2<85%, PaO2<60mmHg and

____ Higher ventilator support needed

PEEP > 8 cm H2O x 12 hrs (10cm x 4 hrs)

FiO2 > 0.8 x 12 hrs (0.9 x 4 hrs)

PIP > 40cm H2O or MAP > 30cmH2O

HFOV

  - ____ Respiratory acidosis (PH < 7.2) under higher ventilator support

  - ____ Air leak syndrome or severe barotraumas

  - ____ Refractory septic shock

_____ Cardiac support:

  - ____ Failure to wean from cardiopulmonary bypass after cardiac surgery

  - ____ Progressive and unremitting low cardiac output state following:

cardiac surgery or myocarditis or cardiomyopathy or septic shock

* low CO = mixed venous saturation < 50%, cardiac index < 2 l/m2/min on

full pharmacological and mechanical support > 2 hrs

  - ____ Prepare for bridge to heart transplantation, heart surgery or VAD.

- ECLS is not recommended:

- ____Major chromosomal anomalies

- ____ Major cardiac or CNS malformations or recent CVA

- ____ Previous chronic lung disease

- ____ Major hemorrhage

- ____ Immunosuppression

- ____ Low certainty of quality of life

- ____ Duration of ventilator support (i.e. FiO2 >0.6 and/or PEEP >5): < 2 years >10 days; 2 - 8 years > 8 days, > 8 years > 6 days (tends to progressive pulmonary fibrosis)

- Major equipment:

1. Pump: centrifugal vs roller-head pump

- Rotation rate: 1200 to 2000 rpm to keep rated flow (CI about 2 L/min/m2)

2. Oxygenator: Hollow fiber oxygenator + heat exchanger+O2 blend+sweep gas flow

3. Heat exchanger: post-oxygenator position to main normothermia (setting < 42℃)

4. Cannula:

Pediatric ECLS cannulae and flows

| BW | <8Kg | 8-12Kg | 12-20Kg | 20-35Kg | >35Kg |
| --- | --- | --- | --- | --- | --- |
| Venous (Fr) | 12-16 | 16-20 | 20-24 | 21-24 | >28 |
| Arterial (Fr) | 10-12 | 12-16 | 14-18 | 16-20 | >20 |
| Rated Flow (l/min) | 1.2 | 1.8 | 1.8-4.5 | 4-5.5 | 5-6.5 |
| Gas flow (l/min) | 2.4 | 4.5 | 7.5 | 10.5 | 13.5 |
| Idle flow (ml/min) | 150 | 300 | 500 | 500-900 | 500-900 |

5. Bladder box: need or not

6. Bridge: come off ECLS: ABV on centrifugal and VBA on roller pump

7. Monitors: ACT, ABG, Saturation (SaO2,SvO2), CBC (Hct), Pressure monitor, Echo etc

Heater monitor, Flow monitor and detector

8. Ports:

proximal to bladder: blood sampling, IV fluid infusion, hemofilter

between pump head and oxygenator: ACT monitor, hemofilter

after oxygenator: pressure monitoring, blood sampling, infusion of platelet

9. Mechanical ventilator: rest setting

FiO2: < 0.6 (0.25~0.4) ; Rate: 6-20 (5-10) ; PIP: 25-35 (20~30) ; PEEP: 8-12 (4-10), Ti: 1sec

- Medications:

1.Heparinization: loading 50-100u/kg; maintenance 5-40u/kg/hr to keep ACT around 200 sec

2.Sedation: midazolam 0.2-0.5 mg/kg/hr ; fentanyl 3-7 mg/kg/hr

3.Broad spectrum antibiotics: to treat existing infections (or to prevent infections)

4.Inotropes: Dopamine, Dobutamine, Primacor, or Epinephrine ect. prn use

5.Fluid: keep I/O balance or negative balance (3/4 to 2/3 maintenance)

6.TPN: higher K, Ca and lower Na supplement ; D10-25W ; protein 2-2.5 g/kg ; lipid 2-3 g/kg/d

7.Routinely consult: respiratory care/infectous disease staff, nephrologist, neurologist etc.

8.Consider IVIG (1 g/kg iv 8 hr) in EV crisis or carditis

9.Consider beta-blocker: propranolol 1-3 mg/kg/day tid in EV crisis and consider esomolol

10.Consider blood exchange transfusion: 2 x blood volume (140-160ml/kg)

11.Check bleeding by surgical intervention: bleeding > half blood volume/day

12.Avoid bleeding: no IM injection, no percutaneous sampling, no new iv setting, iv or urethral catheter left in place.

13.Be careful: avoid blood clot when FFP or WB infusion !!

- Monitors and Laboratory f/u:

1. Vital signs, U/O Q1h and keep BT about 36.5 ~ 37℃

2. Check CBC, Electrolytes (Free Ca, Mg, Phosphorus etc.), Sugar, Urine OB: Q4h ~ Qd prn

3. Check GOT, GPT, albumin, bilirubin, BUN, Cr, CK, CK-MB, fibrinogen, FDP: Qd – Biw prn

4. Continuous pressure monitoring: CVP, PAP, LAP or PWP, BP and ECLS pressure Q1h

venous inlet pressure: CVP (before pump pressure)

before oxygenator pressure:↑pressure (> 300 mmHg) or↑before-after pressure gap

indicate oxygenator malfunction (failure or thrombosis)

after oxygenator pressure: arterial occlusion (tubing kinking)

5. Continuous in-line gas monitoring: SVO2 (Q1h check mixed venous gas in our hospital)

- SVO2 indicate oxygen delivery vs PaO2 indicate oxygenator function

6. Pulse oximeter: SaO2 Q1h

7. ACT monitor: Q1h ~ Q6h

8. CxR, Body weight: Qd

9. Septic work up if needed

10. Echo: heart: Q12h – Qd prn and must rule out cardiac tamponade

brain: Qdprn and must rule out ICH

# Pediatric ECLS parameters and adjustments

| Parameter | Desired | Adjustments |
| --- | --- | --- |
| Hct / Hb | > 40% (12) | Infusion of packed RBC s/p photo |
| Plt count | > 100k | Infusion of platelets s/p photo |
| Blood sugar | 90 - 100 | Infusion of D10-50W |
| ACT | 150 -250 sec | Heparin 5-20 u/kg ; FFP or cryoprecipitation* |
| Fibrinogen | > 150 | FFP or cryo (if < 100) |
| SvO2 | 70 - 80% | ECLS flow, Sweep FiO2 |
| SaO2 | > 90% | ECLS flow, Sweep FiO2 |
| PCO2 | 35-45 mmHg | Sweep gas flow |
| PaO2 | > 65 mmHg | ECLS flow, Sweep FiO2 |

- Pressure monitor, circuit flow and clinical diagnosis:

| Flow | CVP | LAP | BP | Clinical diagnosis |
| --- | --- | --- | --- | --- |
| ( | ( | ( | ( | ECLS overflow or blood/fluid overload,RAA |
| ( | ( | ( | ( | Hypovolemia and/or bleeding |
| ( | ( | ( | ( | Pulmonary hypertension, cardiac tamponade, ECLS malfunction, or ( intrathoracic pressure |
| ( | ( | ( | ( | Sepsis, aortic regurgitation |
| ( | ( | ( | ( | Rv failure |

- Weaning (trial off ECLS):

_____ CxR improved; (PaO2, (VO2, (PaCO2,(Peak airway pressure,(Compliance

_____ LV/RV function improved,(Pulse pressure,(VO2

1. ____ ECLS FiO2: to less 0.4

2. ____ Decrease ECLS flow (10-20ml/min decrements): to idle flow (or 25% of initial flow)

3. ____ Moderate or low mechanical ventilator and consider HFOV

4. ____ Come off ECLS: ABV on centrifugal and VBA on roller pump and decannulation

- Complications and check list:

____ Bleeding: surgical site (27%), GI (6%), intracranial (5%)

____ Infection: 20-30% and consider fungal infection (50%)

____ Pneumothorax, hemothorax, cardiac tamponade

____ Mechanical complications

____ cannula malposition, kinking or tubing rupture

____ clot or air in circuit

____ oxygenator failure

____ pump malfunction

____ heat exchanger malfunction

____ Renal failure: consider hemofiltration and increase pump flow

____ CNS sequela: _____________________________________